Alarm 022

Code 022 refers to the transmission of the tube support and in particular to activation 23S3.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Activation 23S3 absent or not functional LED OK: OFF H | While equipment is switched off insert or replace aboard. |
| Activation powered by an out of range voltage value LED OK: OFF VL | Check board 23S9 and especially fuses, the power supplies and connections from the power transformer, power supplies and line voltage. Replace fuses or board 23S9, restore connections, replace fuses or power transformer 22TR1. |
| Inverted motor connection LED O.K.: OFF LED m.d.: ON L (during quality controls) | Invert motor connection. |
| Motor disconnected LED O.K.: OFF LED O.C.: ON (possible) | Refer to functional diagrams to restore connections. |
| Motor short-circuit LED O.K.: OFF LED O.C.: ON (possible) | Check Motor, the generator/activation tie cable, board 23S9 and activation 23S3. Replace Motor Reducer, tie cable, boards 23S9 and 23S3. |
| Inverted dynamo connection LED O.K.: OFF LED m.d.: ON VVL (during quality controls) | Invert dynamo connection. |
| Disconnected dynamo LED O.K.: OFF LED m.d.: ON L | Refer to functional diagrams to restore connections. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Dinamo short-circuit LED O.K.: OFF LED m.d.: ON VVL | Check generator, the generator/activation tie cable, board 23S9 and activation 23S3. Replace Motor Reducer, tie cable, boards 23S9 and 23S3. |
| Board fault by thermal protection LED O.K.: OFF LED S.T.: ON VL | Use of continuous transmission function. To restore operation switch equipment off, let it cool down and switch it back on. |
| Tracks interruptions on board 23S9 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range VVL | Restore connections or replace board. |
| Interruption in flat-cable 23S9CFC1-25S3CFC1 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range (only during quality controls) | Replace flat-cable. |
| Cold welding on the connectors of board 23S9 and 25S3 Input 25S3-IN11 off Output 25S3-OUT13 off DAC VAN1 out of range VVL (only during quality controls) | Restore joints or replace boards. |
| Input 25S3-IN11 not functional Output 25S3-OUT13 not functional DAC VAN1 out of range VVL | Refer to testing manual to check with Test SW. Replace board I/O 25S3, CPU 25S1 and flat-cable. |

| Events that may cause the error and probability (H=high L=low VL=very low VVL=very very low) | Problem solution |
| --- | --- |
| Power supply failure +24Vcc for input/output board 25S3 VL | Check connections, power supplies on boards 25S3, the power supplies from the service transformer, fuses on transformer and power supplies. Replace board 25S3, restore connections, replace fuses or auxiliaries transformer. |
| Power supply failure +5Vcc or +/-12Vcc for input/output board 25S1 VL | Check connections, power supplies and fuses on boards 25S1, the power supplies from the service transformer, fuses and power supplies on transformer. Replace board 25S1, restore connections, replace fuses or auxiliaries transformer |