Err032: Real Time Clock Data Error 0
Description:
• The real-time clock on the D500 contains invalid data.28 Error Messages
POLYDOROS IT / IT-S COPL-136.840.02.07.02 Siemens AG
08.05 CS PS 212
Page 28 of 52
Medical Solutions
Possible Causes of Error:
• 5 V VCC on the D500 board too low
• Missing configuration data
• D500 defective.
Corrective Measures:
• Adjust the low voltages (Adjustment Instructions)
• D500 reconfiguration (for data: see Operating Instructions); if there is still an error mes-
sage -> replace D500

Err041: Iload Error 0
Description:
• An over-current was detected in the inverter.
Possible Causes of Error:
• Tube unit not correctly warmed up.
• Power too high with unfavorable power line conditions.
• Tube unit defective
• D510 defective
Corrective Measures:
• On HVT/D530, check cable connection X1 (large focus) and X3 (small focus).
• Adjustment: Perform the Static Inverter adjustment
• Adjustment: Perform Learn filament.
• Check the internal line resistance, Ri; for this, see the Startup Instructions.
• Replace tube unit or high voltage tank, but first read out/note the tube data
• Replace the inverter