Alarm 062

Code 062 refers to the transmission of the perpendicular collimator during movement.

Through the position transducer (potentiometer and related circuits) the software determines the incorrect direction of movement.

The activation of this alarm stops the equipment, it disables the power transformer connection relay. All movements are disabled, the angle of tilt display flashes and shows the alarm code while the remaining displays switch off.

| Events that may cause the error | Problem solution |
| --- | --- |
| Tie cable disconnected between board 23S8 and collimator | Reconnect cable. |
| Interruption in tie cable between board 23S8 and collimator | Replace cable. |
| Incorrect connections in tie cable between board 23S8 and collimator | Refer to functional diagrams to restore connections. |
| Tracks interruptions on board 23S10 Input 25S3-IN17 off Output 25S3-OUT23 off Dac VAN6 out of range | Restore connections or replace board. |