4. Optical bench check

◆ Run an initialization and once the instrument is Ready, enter Menu: Services\Diagnostics\Tray .

◆ Check the «Reaction» radio button and press the «Motor Off» button.

◆ Manually remove all cuvette segments from the reaction tray.

◆ Then, press the «Check Home» button.

◆ Enter  Menu: Services\Diagnostics\Spectro\Spectro Settings  ( See “Fig.31: Spectro Settings\Measure Gain”,

page 16).

Fig.31: Spectro Settings\Measure Gain

This tab displays the list of spectrophotometer initialization steps (Fig.31-A) and five subtabs to perform each

initialization step (Fig.31-B). Once a step is performed, it is marked «Done» in the list.

1- Measure Gain

◆ Click the «Measure Gain» tab (Fig.31-B).

◆ Press the «Start» button (Fig.31-C).

◆ When the message «After the gain measurement, the spectrophotometer is initialized.» is displayed, press the

«OK» button.

◆ Then, the following message is displayed: «Verify absence of segment in sector 1.».

Check that there is no cuvette segment into the sector 1 of the reaction tray and press the «OK» button.

◆ Once the step is performed, the gain and the light intens ity measured on air (without  cuvette) are displayed for

each wavelength.

Check that the gains (Fig.31-D) are included between 1 and 95%.

Check that the light intensities (Fig.31-E) are included between 57200 and 65535 (~ 61000).

A

B

C

D E

HORIB A Medi cal conf identi al